	METHOD STATEMENT
	[image: image1.jpg]N







Determinand:
Determination of organochlorine pesticides, synthetic pyrethroids, and selected non-polar organic compounds.

Matrix:

Sample Type: Raw and potable waters.
Principle of Method:

Approximately 200mL of sample is extracted with 5.0+0.2 mL of hexane/ethyl acetate (in a 1:1 ratio mix). The extract is then analysed by gas chromatography using triple quad mass spectrometric detection in MRM mode.

Sampling and Sample Preparation:

Samples are taken in 500mL amber or green glass bottles containing 0.9mL of sodium thiosulfate, 10g L-1, as preservative and are stored at 1-5°C on receipt at the laboratory. Samples should be extracted (or solvent added) within the number of days from sampling as stated in the stability data given below:
	Compound
	Stability – Drinking Water (Days)
	Stability – Surface Water (Days)
	Stability – Ground Water (Days)

	Aldrin
	28
	0
	21

	Chlorbufam
	28
	21
	21

	Chlorothalonil
	21
	0
	7

	Chlorpropham
	28
	21
	21

	Chlorpyrifos-ethyl
	28
	21
	21

	Chlorpyrifos-methyl
	28
	21
	21

	Chlorthal-dimethyl
	28
	21
	21

	Clomazone
	28
	21
	21

	Cypermethrin
	21
	3
	5

	o,p’-DDE
	28
	3
	21

	p,p’-DDE
	28
	3
	21

	o,p’-DDT
	28
	3
	21

	p,p’-DDT
	28
	3
	21

	Deltamethrin
	21
	3
	5

	Diazinon
	28
	21
	21

	Dichlobenil
	28
	21
	21

	Dieldrin
	28
	21
	21

	Diflufenican
	28
	21
	21

	Ethofumesate
	28
	21
	21

	Ethoprophos
	21
	21
	21

	Fenpropimorph
	21
	10
	21

	alpha-HCH
	28
	21
	21

	gamma-HCH
	28
	21
	21

	Heptachlor
	10
	3
	7

	cis-Heptachlor Epoxide
	20
	21
	21

	trans-Heptachlor Epoxide
	28
	21
	21

	Hexachlorobenzene
	28
	3
	21

	Lambda-Cyhalothrin
	28
	0
	21

	Methoxychlor
	28
	14
	21

	Parathion-ethyl
	28
	21
	21

	Pendimethalin
	28
	10
	21

	Pentachlorobenzene
	28
	3
	21

	cis-Permethrin
	28
	3
	21

	trans-Permethrin
	21
	0
	5

	Picolinafen
	28
	10
	21

	Pirimicarb
	28
	21
	21

	Pirimiphos-methyl
	28
	21
	21

	Propham
	28
	21
	21

	Propyzamide
	28
	21
	21

	o,p’-TDE
	28
	21
	21

	p,p’-TDE
	28
	5
	21

	Tecnazene
	28
	21
	21

	Triallate
	28
	21
	21

	Trifluralin
	28
	21
	21


If sample extracts cannot be analysed immediately after preparation, they are stored in a fridge at 5(3oC until analysis can be completed. If the samples show any signs of degradation from the AQC results or blank contamination, then the samples would be re-analysed or re-prepped where feasible.

Interferences
Any compound, which passes through the extraction procedure and that co-elutes with any of the analytes and produces a significant response to the relevant ions being monitored.
Performance of Method:

Range of Application:

The operational range is from the limit of detection, to 150 ng L-1. Samples producing results above this range should be diluted to bring the sample response below that of the top standard.

Limit of Quantification:

Limits of Quantification (LOQ) are given in µg L-1, the highest LOQ has been applied across all instruments. 

	Determinand
	AG2

Limit of Quantification (µg L-1)
	AG3

Limit of Quantification (µg L-1)
	AG5

Limit of Quantification (µg L-1)
	AG6
Limit of Quantification (µg L-1)
	METHOD (Standardised) Limit of Quantification (µg L-1)

	Aldrin
	0.007
	0.005
	0.008
	0.003
	0.008

	Chlorbufam
	0.004
	0.005
	0.007
	0.003
	0.007

	Chlorothalonil
	0.006
	0.003
	0.005
	0.002
	0.006

	Chlorpropham
	0.005
	0.004
	0.003
	0.009
	0.009

	Chlorpyrifos-ethyl
	0.007
	0.008
	0.007
	0.003
	0.008

	Chlorpyrifos-methyl
	0.006
	0.004
	0.008
	0.003
	0.008

	Chlorthal-dimethyl
	0.007
	0.008
	0.008
	0.004
	0.008

	Clomazone
	0.007
	0.005
	0.006
	0.003
	0.007

	Cypermethrin
	0.005
	0.004
	0.008
	0.003
	0.008

	op-DDE
	0.003
	0.005
	0.005
	0.002
	0.005

	pp-DDE 
	0.003
	0.005
	0.004
	0.002
	0.005

	op-DDT
	0.004
	0.004
	0.004
	0.002
	0.004

	pp-DDT
	0.004
	0.004
	0.004
	0.002
	0.004

	Deltamethrin
	0.005
	0.009
	0.006
	0.005
	0.009

	Diazinon
	0.007
	0.003
	0.005
	0.003
	0.007

	Dichlobenil
	0.003
	0.004
	0.004
	0.003
	0.004

	Dieldrin
	0.008
	0.008
	0.009
	0.004
	0.009

	Diflufenican
	0.009
	0.009
	0.006
	0.004
	0.009

	Ethofumesate
	0.007
	0.008
	0.007
	0.003
	0.008

	Ethoprophos
	0.008
	0.006
	0.006
	0.003
	0.008

	Fenpropimorph
	0.010
	0.008
	0.007
	0.005
	0.010

	alpha-HCH 
	0.005
	0.004
	0.006
	0.002
	0.006

	gamma-HCH 
	0.005
	0.004
	0.003
	0.002
	0.005

	Heptachlor
	0.005
	0.004
	0.003
	0.003
	0.005

	cis-Heptachlor Epoxide
	0.007
	0.006
	0.005
	0.002
	0.007

	trans-Heptachlor Epoxide
	0.007
	0.005
	0.007
	0.003
	0.007

	Hexachlorobenzene
	0.005
	0.005
	0.005
	0.002
	0.005

	Lambda-Cyhalothrin
	0.006
	0.005
	0.006
	0.003
	0.006

	Methoxychlor
	0.006
	0.004
	0.004
	0.002
	0.006

	Parathion
	0.006
	0.010
	0.012
	0.003
	0.012

	Pendimethalin
	0.009
	0.009
	0.009
	0.004
	0.009

	Pentachlorobenzene
	0.006
	0.003
	0.004
	0.002
	0.006

	cis-Permethrin
	0.004
	0.005
	0.006
	0.004
	0.006

	trans-Permethrin
	0.004
	0.005
	0.004
	0.002
	0.005

	Picolinafen
	0.009
	0.007
	0.006
	0.004
	0.009

	Pirimicarb
	0.014
	0.010
	0.016
	0.006
	0.016

	Pirimiphos-methyl
	0.008
	0.005
	0.005
	0.003
	0.008

	Propham
	0.004
	0.004
	0.005
	0.002
	0.005

	Propyzamide
	0.009
	0.008
	0.007
	0.006
	0.009

	op-TDE
	0.004
	0.004
	0.005
	0.002
	0.005

	pp-TDE
	0.003
	0.003
	0.005
	0.002
	0.005

	Tecnazene
	0.008
	0.005
	0.004
	0.002
	0.008

	Triallate
	0.006
	0.004
	0.006
	0.002
	0.006

	Trifluralin
	0.007
	0.005
	0.004
	0.002
	0.007


Recoveries of Compounds and Uncertainty of measurement:

Instrument, WAKGCQQQ02, AG2:
	Determinand
	Low Standard 30ng/L Spike
	High Standard 100ng/L Spike
	Elvington Treated Water (Hard Water) PCV Spike
	Uncertainty

	
	% Recovery
	% RSD
	% Recovery
	% RSD
	% Recovery
	% RSD
	

	Aldrin
	98.78%
	4.78%
	101.20%
	7.37%
	99.39%
	3.35%
	± 9.86 %

	Chlorbufam
	97.72%
	6.79%
	98.28%
	4.17%
	95.69%
	4.13%
	± 11.67 %

	Chlorothalonil
	92.46%
	6.06%
	99.44%
	6.72%
	106.90%
	7.20%
	± 19.08 %

	Chlorpropham
	97.72%
	5.75%
	100.69%
	3.14%
	99.65%
	3.51%
	± 7.71 %

	Chlorpyrifos-ethyl
	99.39%
	6.79%
	100.24%
	5.35%
	97.73%
	3.17%
	± 12.62 %

	Chlorpyrifos-methyl
	97.81%
	5.30%
	100.19%
	3.03%
	100.38%
	3.82%
	± 9.98 %

	Chlorthal-dimethyl 
	97.67%
	2.82%
	100.36%
	8.24%
	98.91%
	2.61%
	± 9.20 %

	Clomazone
	96.00%
	7.55%
	98.47%
	2.82%
	95.48%
	2.64%
	± 12.40 %

	Cypermethrin
	96.18%
	3.67%
	100.64%
	6.56%
	97.55%
	2.94%
	± 13.29 %

	Deltamethrin
	97.12%
	4.60%
	101.00%
	5.21%
	98.83%
	5.09%
	± 10.07 %

	Diazinon
	97.34%
	4.61%
	101.18%
	7.69%
	99.83%
	2.86%
	± 8.90 %

	Dichlobenil
	106.86%
	7.13%
	102.64%
	6.40%
	99.61%
	5.95%
	± 11.41 %

	Dieldrin
	95.65%
	3.69%
	100.30%
	10.61%
	95.66%
	4.95%
	± 11.29 %

	Diflufenican
	96.16%
	4.04%
	99.82%
	6.96%
	98.23%
	3.00%
	± 10.84 %

	Ethofumesate
	98.30%
	7.74%
	99.76%
	6.26%
	99.78%
	6.26%
	± 15.67 %

	Ethoprophos
	98.02%
	6.79%
	95.40%
	5.05%
	91.64%
	4.09%
	± 17.72 %

	Fenpropimorph
	99.24%
	6.52%
	101.01%
	5.11%
	98.58%
	4.62%
	± 19.90 %

	HCH-alpha
	105.25%
	2.68%
	102.68%
	4.78%
	100.72%
	2.06%
	± 8.28 %

	HCH-gamma
	101.02%
	2.75%
	101.87%
	4.56%
	99.98%
	2.05%
	± 7.26 %

	Heptachlor
	93.87%
	7.23%
	99.75%
	6.05%
	100.98%
	4.75%
	± 9.68 %

	cis-Heptachlor Epoxide
	96.76%
	4.34%
	99.60%
	9.06%
	98.59%
	3.83%
	± 9.44 %

	trans-Heptachlor Epoxide
	99.69%
	3.46%
	100.35%
	7.44%
	99.27%
	3.16%
	± 12.41 %

	Hexachlorobenzene
	105.42%
	3.44%
	103.20%
	5.23%
	100.98%
	3.00%
	± 8.54 %

	lambda-Cyhalothrin
	99.19%
	9.91%
	102.63%
	11.81%
	102.04%
	9.67%
	± 13.73 %

	Methoxychlor
	98.83%
	12.73%
	101.10%
	8.06%
	93.51%
	4.71%
	± 22.74 %

	op-DDE
	99.92%
	2.34%
	101.37%
	5.33%
	99.71%
	1.18%
	± 6.25 %

	op-DDT
	96.06%
	1.66%
	99.94%
	5.99%
	98.76%
	1.92%
	± 9.00 %

	op-TDE
	99.19%
	3.74%
	100.56%
	7.40%
	95.12%
	2.77%
	± 13.17 %

	Parathion-ethyl 
	98.34%
	7.45%
	99.59%
	6.51%
	98.17%
	3.63%
	± 10.99 %

	Pendimethalin
	100.16%
	11.89%
	98.88%
	5.06%
	100.38%
	6.33%
	± 16.08 %

	Pentachlorobenzene 
	104.83%
	5.88%
	102.72%
	5.71%
	100.47%
	4.60%
	± 9.78 %

	cis-Permethrin
	96.89%
	4.87%
	100.74%
	6.79%
	99.14%
	3.30%
	± 9.73 %

	trans-Permethrin
	96.68%
	4.74%
	100.72%
	6.39%
	98.40%
	3.32%
	± 9.73 %

	Picolinafen
	97.88%
	6.14%
	99.46%
	3.13%
	99.29%
	4.50%
	± 10.54 %

	Pirimicarb
	99.83%
	9.16%
	100.04%
	6.09%
	100.21%
	9.52%
	± 21.71 %

	Pirimiphos-methyl
	94.46%
	6.09%
	97.66%
	4.29%
	96.15%
	4.17%
	± 12.89 %

	pp-DDE
	98.61%
	1.60%
	100.93%
	5.86%
	98.73%
	4.84%
	± 7.22 %

	pp-DDT
	94.69%
	5.89%
	99.19%
	6.12%
	99.71%
	2.89%
	± 10.26 %

	pp-TDE
	95.27%
	6.22%
	100.33%
	7.48%
	94.27%
	4.53%
	± 17.92 %

	Propham
	100.14%
	4.26%
	100.24%
	3.19%
	98.99%
	2.80%
	± 7.89 %

	Propyzamide
	96.31%
	6.84%
	96.98%
	4.37%
	93.90%
	4.96%
	± 15.46 %

	Tecnazene
	103.03%
	3.62%
	100.54%
	6.28%
	98.25%
	6.42%
	± 9.57 %

	Triallate
	95.01%
	3.40%
	99.53%
	8.71%
	97.89%
	3.49%
	± 10.42 %

	Trifluralin
	95.55%
	2.73%
	100.38%
	6.06%
	99.65%
	2.34%
	± 8.79 %


Instrument, WAKGCQQQ03, AG3:

	Determinand
	Low Standard 30ng/L Spike
	High Standard 100ng/L Spike
	Elvington Treated Water (Hard Water) PCV Spike
	Uncertainty

	
	% Recovery
	RSD
	% Recovery
	RSD
	% Recovery
	RSD
	

	Aldrin
	101.56%
	4.57%
	101.84%
	7.25%
	98.01%
	4.65%
	± 20.90 %

	Chlorbufam
	96.90%
	8.00%
	98.67%
	3.43%
	96.03%
	3.40%
	± 14.74 %

	Chlorothalonil
	96.87%
	7.89%
	98.79%
	8.16%
	103.92%
	10.55%
	± 11.18 %

	Chlorpropham
	96.19%
	5.80%
	100.18%
	3.75%
	100.10%
	2.27%
	± 9.28 %

	Chlorpyrifos-ethyl
	96.05%
	5.99%
	97.28%
	5.10%
	98.74%
	5.00%
	± 14.91 %

	Chlorpyrifos-methyl
	98.88%
	5.23%
	100.45%
	4.32%
	99.66%
	4.24%
	± 10.16 %

	Chlorthal-dimethyl 
	101.24%
	8.47%
	100.10%
	8.41%
	100.89%
	8.07%
	± 12.61 %

	Clomazone
	98.66%
	8.12%
	99.06%
	4.36%
	96.61%
	4.49%
	± 12.85 %

	Cypermethrin
	93.46%
	4.11%
	98.63%
	3.04%
	97.77%
	2.45%
	± 18.33 %

	Deltamethrin
	89.83%
	4.96%
	98.73%
	3.95%
	97.12%
	2.42%
	± 20.59 %

	Diazinon
	95.69%
	6.41%
	98.71%
	2.71%
	99.86%
	3.53%
	± 15.21 %

	Dichlobenil
	103.60%
	9.74%
	101.27%
	6.56%
	101.11%
	7.40%
	± 20.61 %

	Dieldrin
	98.86%
	4.05%
	101.99%
	2.73%
	99.08%
	3.73%
	± 15.81 %

	Diflufenican
	95.44%
	2.80%
	99.15%
	1.95%
	99.39%
	2.52%
	± 14.90 %

	Ethofumesate
	100.27%
	8.97%
	99.27%
	3.57%
	100.31%
	6.07%
	± 17.94 %

	Ethoprophos
	93.10%
	7.74%
	95.36%
	5.12%
	93.75%
	6.55%
	± 18.68 %

	Fenpropimorph
	96.85%
	9.31%
	100.70%
	7.48%
	99.93%
	5.43%
	± 13.44 %

	HCH-alpha
	103.97%
	8.36%
	102.45%
	7.83%
	102.87%
	8.16%
	± 17.01 %

	HCH-gamma
	100.78%
	1.52%
	100.46%
	1.38%
	99.97%
	1.55%
	± 12.31 %

	Heptachlor
	99.73%
	8.49%
	99.75%
	6.53%
	102.95%
	9.03%
	± 24.06 %

	cis-Heptachlor Epoxide
	97.29%
	6.06%
	99.74%
	6.14%
	97.45%
	3.49%
	± 15.58 %

	trans-Heptachlor Epoxide
	95.57%
	5.23%
	99.51%
	4.99%
	96.22%
	4.62%
	± 15.71 %

	Hexachlorobenzene
	105.49%
	6.98%
	103.45%
	7.59%
	102.34%
	6.93%
	± 14.55 %

	lambda-Cyhalothrin
	96.22%
	8.42%
	99.46%
	5.70%
	99.40%
	11.14%
	± 17.70 %

	Methoxychlor
	96.04%
	7.02%
	98.60%
	6.13%
	95.96%
	5.62%
	± 21.05 %

	op-DDE
	102.72%
	2.81%
	99.98%
	5.53%
	99.79%
	2.77%
	± 13.94 %

	op-DDT
	98.92%
	1.93%
	100.47%
	1.87%
	99.43%
	2.04%
	± 13.39 %

	op-TDE
	103.49%
	2.10%
	101.51%
	2.16%
	98.69%
	2.36%
	± 13.51 %

	Parathion-ethyl 
	94.39%
	5.21%
	99.71%
	2.91%
	99.71%
	3.24%
	± 14.84 %

	Pendimethalin
	96.17%
	7.76%
	98.53%
	5.61%
	97.79%
	4.62%
	± 16.79 %

	Pentachlorobenzene 
	102.58%
	8.93%
	103.27%
	6.87%
	102.35%
	7.18%
	± 20.66 %

	cis-Permethrin
	94.52%
	6.88%
	98.86%
	4.20%
	98.68%
	3.64%
	± 14.72 %

	trans-Permethrin
	93.68%
	7.66%
	99.34%
	3.76%
	99.84%
	3.84%
	± 17.87 %

	Picolinafen
	96.83%
	6.25%
	99.24%
	4.18%
	98.92%
	6.18%
	± 14.92 %

	Pirimicarb
	103.73%
	9.65%
	100.68%
	9.84%
	100.20%
	6.73%
	± 17.18 %

	Pirimiphos-methyl
	96.78%
	6.57%
	99.00%
	3.77%
	97.49%
	3.87%
	± 11.61 %

	pp-DDE
	101.37%
	2.20%
	100.79%
	2.32%
	99.83%
	2.30%
	± 13.05 %

	pp-DDT
	94.48%
	2.59%
	99.09%
	1.57%
	100.79%
	2.39%
	± 15.44 %

	pp-TDE
	100.19%
	4.24%
	101.17%
	3.47%
	97.75%
	3.77%
	± 16.52 %

	Propham
	98.29%
	5.48%
	100.05%
	2.68%
	99.05%
	3.76%
	± 8.84 %

	Propyzamide
	99.51%
	8.35%
	97.88%
	5.32%
	94.69%
	6.27%
	± 16.86 %

	Tecnazene
	99.06%
	8.11%
	101.36%
	6.56%
	103.08%
	7.70%
	± 17.39 %

	Triallate
	97.83%
	3.41%
	100.46%
	4.55%
	102.22%
	4.99%
	± 12.65 %

	Trifluralin
	93.02%
	2.99%
	99.21%
	2.90%
	99.29%
	2.41%
	± 17.54 %


Instrument, WAKGCQQQ05, AG5:

	Determinand
	Low Standard 30ng/L Spike
	High Standard 100ng/L Spike
	Elvington Treated Water (Hard Water) PCV Spike
	Uncertainty

	
	% Recovery
	RSD
	% Recovery
	RSD
	% Recovery
	RSD
	

	Aldrin
	97.38%
	4.92%
	101.61%
	2.89%
	98.80%
	4.55%
	±15.75%

	Chlorbufam
	94.43%
	6.92%
	94.91%
	4.25%
	95.02%
	5.12%
	± 15.57%

	Chlorothalonil
	91.89%
	7.86%
	99.96%
	3.42%
	100.99%
	2.69%
	±17.50%

	Chlorpropham
	95.99%
	5.16%
	98.93%
	4.41%
	100.39%
	4.08%
	± 12.25%

	Chlorpyrifos-ethyl
	95.54%
	8.74%
	98.44%
	3.34%
	98.28%
	3.91%
	± 10.68%

	Chlorpyrifos-methyl
	95.18%
	7.17%
	99.46%
	3.47%
	97.71%
	3.67%
	± 10.32%

	Chlorthal-dimethyl 
	99.16%
	4.59%
	100.87%
	2.74%
	101.40%
	2.82%
	±11.15%

	Clomazone
	97.63%
	6.50%
	97.31%
	4.13%
	98.03%
	5.24%
	± 12.27%

	Cypermethrin
	99.21%
	8.30%
	101.24%
	3.46%
	100.75%
	3.32%
	±9.17%

	Deltamethrin
	95.45%
	5.65%
	98.04%
	3.18%
	99.19%
	2.68%
	±13.23%

	Dichlobenil
	101.50%
	2.60%
	101.51%
	3.74%
	99.44%
	4.22%
	±10.54%

	Diazinon
	96.64%
	2.98%
	99.67%
	2.56%
	100.90%
	3.46%
	±8.26%

	Dieldrin
	100.04%
	3.01%
	101.19%
	2.23%
	99.80%
	2.88%
	±12.11%

	Diflufenican
	97.30%
	2.61%
	100.35%
	1.62%
	100.33%
	2.77%
	±8.49%

	Ethofumesate
	97.45%
	6.17%
	97.97%
	4.30%
	100.05%
	4.97%
	± 15.35%

	Ethoprophos
	92.63%
	7.16%
	94.29%
	3.85%
	92.97%
	5.50%
	± 19.47%

	Fenpropimorph
	98.45%
	7.78%
	101.31%
	3.50%
	101.17%
	4.40%
	± 13.11%

	HCH-alpha
	101.14%
	3.13%
	101.90%
	2.21%
	101.84%
	3.41%
	±8.45%

	HCH-gamma
	100.44%
	2.47%
	102.12%
	1.90%
	101.26%
	3.04%
	±7.93%

	Heptachlor
	96.82%
	3.81%
	101.42%
	1.83%
	99.41%
	3.21%
	±9.51%

	cis-Heptachlor Epoxide
	101.31%
	3.57%
	101.66%
	4.02%
	101.75%
	5.38%
	±13.06%

	trans-Heptachlor Epoxide
	99.05%
	4.31%
	101.50%
	3.53%
	99.77%
	5.66%
	±12.45%

	Hexachlorobenzene
	101.50%
	3.08%
	103.64%
	2.31%
	102.60%
	2.87%
	±9.22%

	lambda-Cyhalothrin
	100.97%
	8.97%
	99.87%
	4.53%
	98.98%
	4.82%
	±14.79%

	Methoxychlor
	99.54%
	6.59%
	101.59%
	2.89%
	101.55%
	2.41%
	±8.57%

	op-DDE
	100.54%
	2.76%
	101.97%
	2.66%
	101.71%
	3.63%
	±8.17%

	op-DDT
	99.83%
	1.89%
	100.63%
	1.22%
	99.71%
	2.64%
	±7.73%

	op-TDE
	97.04%
	3.00%
	100.39%
	1.91%
	100.36%
	3.24%
	±9.72%

	Parathion-ethyl 
	95.11%
	7.35%
	98.04%
	3.48%
	98.98%
	3.41%
	±12.02%

	Pendimethalin
	99.25%
	8.35%
	99.34%
	5.28%
	98.04%
	6.81%
	± 18.95%

	Pentachlorobenzene 
	102.33%
	2.72%
	103.88%
	3.45%
	103.24%
	3.03%
	±10.78%

	cis-Permethrin
	98.82%
	7.24%
	102.24%
	4.70%
	101.36%
	4.38%
	±12.80%

	trans-Permethrin
	99.08%
	5.01%
	101.83%
	3.35%
	101.10%
	3.93%
	±12.09%

	Picolinafen
	97.02%
	7.33%
	99.95%
	3.68%
	99.25%
	3.85%
	± 10.56%

	Pirimicarb
	103.97%
	9.39%
	101.03%
	5.33%
	98.90%
	4.81%
	± 16.59%

	Pirimiphos-methyl
	94.38%
	8.91%
	97.82%
	3.81%
	98.42%
	3.96%
	± 12.26%

	pp-DDE
	98.45%
	3.00%
	100.42%
	2.30%
	100.89%
	3.02%
	±8.44%

	pp-DDT
	94.34%
	2.64%
	98.92%
	2.00%
	99.15%
	2.69%
	±12.51%

	pp-TDE
	96.95%
	2.15%
	100.05%
	2.99%
	99.00%
	2.33%
	±11.82%

	Propham
	98.68%
	4.26%
	99.68%
	3.82%
	100.15%
	3.48%
	± 11.74%

	Propyzamide
	97.63%
	7.77%
	97.67%
	4.59%
	97.77%
	5.82%
	± 14.33%

	Tecnazene
	97.81%
	2.69%
	101.15%
	2.60%
	100.83%
	3.23%
	±10.69%

	Triallate
	97.99%
	2.61%
	100.66%
	2.01%
	101.65%
	2.73%
	±8.41%

	Trifluralin
	100.89%
	6.12%
	100.85%
	1.92%
	101.50%
	2.65%
	±7.28%


	Determinand
	Low Standard 30ng/L Spike
	High Standard 100ng/L Spike
	Elvington Treated Water (Hard Water) PCV Spike
	Uncertainty

	
	% Recovery
	RSD
	% Recovery
	RSD
	% Recovery
	RSD
	

	Aldrin
	99.17%
	2.10%
	100.52%
	1.47%
	100.90%
	1.70%
	± 9.63%

	Chlorbufam
	94.55%
	3.36%
	97.95%
	2.94%
	101.15%
	4.01%
	± 13.23%

	Chlorothalonil
	93.64%
	2.24%
	98.14%
	2.03%
	98.11%
	2.61%
	± 12.79%

	Chlorpropham
	96.03%
	2.83%
	97.29%
	3.02%
	99.31%
	3.04%
	± 10.99%

	Chlorpyrifos-ethyl
	97.58%
	3.62%
	99.97%
	3.12%
	99.22%
	2.17%
	± 9.58%

	Chlorpyrifos-methyl
	95.33%
	2.86%
	99.68%
	1.40%
	99.44%
	1.51%
	± 9.74%

	Chlorthal-dimethyl 
	95.26%
	3.30%
	98.87%
	2.37%
	98.54%
	1.99%
	± 12.36%

	Clomazone
	94.26%
	2.91%
	97.55%
	2.75%
	100.85%
	2.78%
	± 12.75%

	Cypermethrin
	93.15%
	2.35%
	97.82%
	2.46%
	99.35%
	2.96%
	± 12.22%

	op-DDE
	97.41%
	1.58%
	99.37%
	1.23%
	99.91%
	1.31%
	± 8.46%

	pp-DDE
	96.46%
	1.25%
	99.32%
	1.23%
	99.93%
	1.25%
	± 8.13%

	op-DDT
	94.49%
	1.69%
	98.64%
	1.15%
	99.61%
	1.14%
	± 9.38%

	pp-DDT
	92.45%
	1.44%
	98.02%
	0.99%
	98.64%
	1.27%
	± 10.81%

	Deltamethrin
	90.19%
	2.98%
	98.03%
	2.36%
	98.26%
	3.00%
	± 15.67%

	Diazinon
	92.01%
	2.69%
	97.92%
	1.85%
	97.75%
	2.50%
	± 14.00%

	Dichlobenil
	101.19%
	3.16%
	100.37%
	3.87%
	99.95%
	4.50%
	± 11.53%

	Dieldrin
	96.96%
	1.75%
	99.67%
	2.21%
	99.15%
	2.46%
	± 9.71%

	Diflufenican
	95.54%
	2.81%
	99.50%
	1.34%
	99.80%
	1.63%
	± 9.77%

	Ethofumesate
	94.19%
	2.91%
	98.64%
	2.14%
	99.26%
	2.65%
	± 12.09%

	Ethoprophos
	94.17%
	3.58%
	98.49%
	3.29%
	102.36%
	3.91%
	± 13.20%

	Fenpropimorph
	89.57%
	4.94%
	96.46%
	2.49%
	96.17%
	3.12%
	± 23.02%

	HCH-alpha
	100.10%
	1.96%
	100.30%
	1.69%
	100.02%
	1.91%
	± 8.84%

	HCH-gamma
	100.59%
	1.71%
	100.68%
	1.43%
	100.26%
	2.03%
	± 9.96%

	Heptachlor
	95.54%
	2.56%
	99.57%
	1.71%
	97.14%
	2.18%
	± 11.23%

	cis-Heptachlor Epoxide
	98.18%
	2.08%
	99.69%
	1.70%
	99.63%
	2.08%
	± 9.04%

	trans-Heptachlor Epoxide
	98.72%
	1.93%
	99.67%
	2.03%
	100.54%
	2.03%
	± 9.14%

	Hexachlorobenzene
	100.47%
	2.16%
	100.87%
	1.45%
	100.28%
	2.49%
	± 9.48%

	Lambda-Cyhalothrin
	92.70%
	10.71%
	98.27%
	4.48%
	98.80%
	4.65%
	± 10.44%

	Methoxychlor
	86.76%
	2.29%
	96.95%
	1.58%
	96.37%
	1.41%
	± 20.70%

	Parathion-ethyl 
	94.26%
	3.28%
	99.83%
	2.76%
	99.64%
	2.42%
	± 14.33%

	Pendimethalin
	92.58%
	6.35%
	97.08%
	3.18%
	98.52%
	2.64%
	± 14.82%

	Pentachlorobenzene 
	101.06%
	1.89%
	101.37%
	1.50%
	100.37%
	2.65%
	± 10.65%

	cis-Permethrin
	90.76%
	10.96%
	98.69%
	3.94%
	99.94%
	5.23%
	± 21.15%

	trans-Permethrin
	88.29%
	8.32%
	97.71%
	3.80%
	100.05%
	3.69%
	± 17.72%

	Picolinafen
	93.35%
	2.72%
	98.22%
	2.48%
	98.98%
	1.72%
	± 12.45%

	Pirimicarb
	97.80%
	4.17%
	99.44%±
	3.66%
	99.62%
	3.36%
	± 13.61%

	Pirimiphos-methyl
	91.99%
	3.08%
	98.13%
	1.96%
	96.96%
	2.59%
	± 15.87%

	Propham
	98.69%
	2.53%
	98.94%
	2.97%
	101.54%
	2.60%
	± 9.30%

	Propyzamide
	92.06%
	4.82%
	95.96%
	3.50%
	100.47%
	2.59%
	± 13.54%

	op-TDE
	96.40%
	1.60%
	99.35%
	1.44%
	99.65%
	1.81%
	± 9.07%

	pp-TDE
	94.10%
	4.10%
	98.44%
	3.32%
	100.01%
	4.14%
	± 14.56%

	Tecnazene
	99.60%
	2.45%
	100.54%
	1.50%
	100.42%
	1.50%
	± 9.21%

	Triallate
	96.51%
	2.56%
	99.54%
	1.82%
	100.24%
	2.25%
	± 10.99%

	Trifluralin
	95.32%
	2.10%
	99.20%
	1.39%
	99.48%
	1.21%
	± 10.71%


References:

In house method
	WPC50
	Last Updated: 08/06/26
	Page 4 of 6



[image: image1.jpg]